## Slide 1: How to evaluate/treat and random/systematic errors in EPMA
John J. Donovan, CAMCOR, University of Oregon

## Slide 2
Our random errors are described by simple photon counting statistics
Since we cannot count less than zero photons, these statistics are best described by Poisson statistics…
However, these calculations involve an exponential and a factorial which are cumbersome to handle, so we therefore choose to describe our photon counting statistics using Gaussian (also known as normal) statistics…
Which is described by simply taking the square root of the raw count variance…
Standard deviation =

## Slide 3
K-ratio vs. acquisition #
Errors in precision or sensitivity can always be improved by:
averaging more measurements
or counting longer!
Sensitivity vs. acquisition #
2x time
SQRT sensitivity
4x time
2x
sensitivity

## Slide 4
We see the same trends by increasing beam currents:
K-ratio vs. acquisition #
Sensitivity vs. acquisition #

## Slide 5
The “takeaway” from all this?
Double your counting time
and double your beam current
to double your sensitivity!

## Slide 6
We should use std dev for single data points, std err for averages

## Slide 7
Exception for t-test: use Std Deviation!

## Slide 8
Systematic (accuracy) errors
incorrect values for composition of standards
instrumental instability (temperature effect on crystal 2d, stage Z drift as it heats up)
inappropriate matrix corrections, “unanalyzed” (missing or excess) elements
poor electrical ground of standard and/or unknown
carbon coat variation (2 effects- one for soft x-ray, another for hard x-rays)
beam change/damage to unknown (e.g. TDI effects for Na, K in glass)
difference in peak shape/position for light elements (standard vs unknown)
on-peak or background (off-peak) spectral interferences
pulse height depression on standards, pulse height gain on traces, dead time calibrations
fluorescence across observed phase boundaries (e.g. diffusion couple)
inaccurate calibration of accelerating voltage (use Duane-Hunt limit)

## Slide 9
Intensity
Concentration
Background (subtraction)
Primary standard (multiplicative)
Dominant effects:
Standard
Matrix
Deadtime
Dominant effects:
Background
Interferences
0
100
0

## Slide 10
Spectral Interference Corrections

## Slide 11
Errors due to “unanalyzed” elements!
Examples:
Excess oxygen when ferric iron is present
Carbon in carbonates
Boron, lithium, etc.
When water or hydroxyl are present in minerals and glasses
Fournelle et al., 2020
Roman et al., 2006

## Slide 12
“Trace” Ti in quartz adjacent to rutile (not chemical diffusion!)
“Trace” Co in copper adjacent to
cobalt (not chemical diffusion!)
Borisova et al., Secondary Fluorescence effects in microbeam analysis, Chemical Geology, 490, 22-29, (2018)

## Slide 13
Accuracy Check!
Trace Accuracy: How accurately can you measure zero?
“Holes” in the
continuum!
Requires a “blank” correction for high accuracy traces!
Ti ka in Quartz

## Slide 14
Sample/standard Error: Beam Sensitivity, cont’d
SRM K-1718 20 nA, 0 um, 1 sec intervals, Na Ka
SRM K-1718 20 nA, 0 um, 1 sec intervals, Si Ka
Y axis: Log Scale

## Slide 15
Time Dependent Intensity (TDI ) Corrections in Quantitative X-ray mapping
Without TDI correction
With TDI correction